## Cover
Agilent 871xE series and ENA 6 series command comparison table
This spread sheet provides a comparison table for Agilent 871xE series scpi commands and those of E5061A/62A/61B
Please refer to the next sheet for detail.
Notices
The information contained in this file is subject to change without notice.
This file contains proprietary information that is protected by copyright. All rights
are reserved. No part of this document may be photocopied, reproduced, or translated to
another language without the prior written consent of the Agilent Technologies.
Exclusive Remedies
The remedies provided herein are buyer’s sole and exclusive remedies. Agilent
Technologies shall not be liable for any direct, indirect, special, incidental, or
consequential damages, whether based on contract, tort, or any other legal theory.
History
Revision 1.0 January 2004
Revision 1.1 August 2012
Status summary:
Compatibility Level | Description
A+ | Completely compatible
A | 1 equivalent command
B | Multiple candidate commands type 1 more than 1 commands are necessary
C | Multiple candidate commands type 2 some features are not supported
D | No equivalent command

## Command list
Code Conversion table Rev 1.1 | A+: 51, A: 32, B: 77 , C: 58 , D: 209 as of 18 June 2012
Last update 18 June 2012 | Status : A+: completely compatible A: 1 equivalent command B: several commands are required C: several candidate, needs re-write or considaration D: no equivalent command
871x |  | ENA 6 series
8712/8714E series scpi commands | Discription | ENA 6 series compatibility status | conversion? | E5061/62A series Correspond commands | E5061B series Correspond commands | Tips? | Remarks
CALCulate[1¦2]:DATA? | Queries the formatted data trace. (Same as TRAC? CH<1¦2>FDATA) | B | T | :CALCulate[1-4][:SELected]:DATA:FDATa | :CALCulate[1-4][:SELected]:DATA:FDATa | Data | When plural traces are used in eligible channel, need to select an eligible trace by ":CALC[1-4]:PAR[1-4]:SEL" first.
CALCulate[1¦2]:FORMat <char> | Selects the display format for measurement data | B | T | :CALCulate[1-4][:SELected]:FORMat | :CALCulate[1-4][:SELected]:FORMat |  | To achieve "MIMPedance", Need to use "CALC[1-4]:PCONV:FUNC:ZREFlection" and ":CALC[1-4]:CONV ON"
CALCulate[1¦2]:FORMat:UNIT:MLIN <char> | Selects linear magnitude units for Y-axis display. (W¦MW¦UW¦V¦MV¦UV) | D | F
CALCulate[1¦2]:FORMat:UNIT:MLOG <char> | Selects log magnitude units for Y-axis display. (DBW¦DBM¦DBUW¦DBV¦DBMV¦DBUV) | D | F
CALCulate[1¦2]:GDAPerture:APERture <num> | Sets the group delay aperture as aperture/span. | B | T | :SENS{1-4}:FREQ:SPAN/STAR/STOP/CNT | :SENS{1-4}:FREQ:SPAN/STAR/STOP/CNT |  | Direct conversion is impossible. Use frequency span setting
CALCulate[1¦2]:GDAPerture:SPAN <num> | Specifies the group delay aperture in Hz. | B | T | :SENS{1-4}:FREQ:SPAN/STAR/STOP/CNT | :SENS{1-4}:FREQ:SPAN/STAR/STOP/CNT |  | Direct conversion is impossible. Use frequency span setting
CALCulate[1¦2]:LIMit:DISPlay <ON¦OFF> | Turns display of limit lines on/off. | A+ | T | :CALCulate[1-4][:SELected]:LIMit:DISPlay[:STATe] | :CALCulate[1-4][:SELected]:LIMit:DISPlay[:STATe] | Limit
CALCulate[1¦2]:LIMit:MARKer:FLATness:MAXimum <num> | Sets the maximum value for a flatness limit test. | D | F |  |  |  | ENA-L has "marker flatness" function itself, althouhg marker limit test is not supported.
CALCulate[1¦2]:LIMit:MARKer:FLATness:MINimum <num> | Sets the minimum value for a flatness marker limit test. | D | F |  |  |  | ENA-L has "marker flatness" function itself, althouhg marker limit test is not supported.
CALCulate[1¦2]:LIMit:MARKer:FLATness:STATe <ON¦OFF> | Turns flatness marker limit test on/off. | D | F |  |  |  | ENA-L has "marker flatness" function itself, althouhg marker limit test is not supported.
CALCulate[1¦2]:LIMit:MARKer:FREQuency:MAXimum <num> | Sets the maximum value for delta frequency marker limit test. | D | F
CALCulate[1¦2]:LIMit:MARKer:FREQuency:MINimum <num> | Sets the minimum value for delta frequency marker limit test. | D | F
CALCulate[1¦2]:LIMit:MARKer:FREQuency [:STATe] <ON¦OFF> | Turns delta frequency marker limit testing on or off. | D | F
CALCulate[1¦2]:LIMit:MARKer:STATistic:MEAN:MAXimum <num> | Sets the maximum value for a statistic mean limit test. | D | F
CALCulate[1¦2]:LIMit:MARKer:STATistic:MEAN:MINimum <num> | Sets the minimum value for a statistic mean limit test. | D | F
CALCulate[1¦2]:LIMit:MARKer:STATistic:MEAN:STATe <ON¦OFF> | Turn statistic mean marker limit test on/off. | D | F
CALCulate[1¦2]:LIMit:MARKer:STATistic:PEAK:MAXimum <num> | Sets the maximum value for a statistic peak-to-peak limit test. | D | F
CALCulate[1¦2]:LIMit:MARKer:STATistic:PEAK:MINimum <num> | Sets the minimum value for a statistic peak-to-peak limit test. | D | F
CALCulate[1¦2]:LIMit:MARKer:STATistic:PEAK:STATe <ON¦OFF> | Turns statistic peak-to-peak marker limit test on/off. | D | F
CALCulate[1¦2]:LIMit:MARKer:TILT:MAXimum <num> | Sets the maximum value for delta amplitude marker limit test. | D | F
CALCulate[1¦2]:LIMit:MARKer:TILT:MINimum <num> | Sets the minimum value for delta amplitude marker limit test. | D | F
CALCulate[1¦2]:LIMit:MARKer:TILT [:STATe] <ON¦OFF> | Turns delta amplitude marker limit testing on or off. | D | F
CALCulate[1¦2]:LIMit:SEGMent[1¦2¦12]:AMPLitude:STARt <num> | Sets the Begin Limit for the specified limit segment. | B | T | :CALCulate[1-4][:SELected]:LIMit:DATA | :CALCulate[1-4][:SELected]:LIMit:DATA | Limit
CALCulate[1¦2]:LIMit:SEGMent[1¦2¦12]:AMPLitude:STOP <num> | Sets the End Limit for the specified limit segment. | B | T | :CALCulate[1-4][:SELected]:LIMit:DATA | :CALCulate[1-4][:SELected]:LIMit:DATA | Limit
CALCulate[1¦2]:LIMit:SEGMent:AOFF | Turns off all limit segments for a given channel | B | T | :CALCulate[1-4][:SELected]:LIMit:DATA | :CALCulate[1-4][:SELected]:LIMit:DATA | Limit
CALCulate[1¦2]:LIMit:SEGMent[1¦2¦12]:DISTance:STARt <num> | Sets the Begin Distance for the specified limit segment. | B | T | :CALCulate[1-4][:SELected]:LIMit:DATA | :CALCulate[1-4][:SELected]:LIMit:DATA | Limit
CALCulate[1¦2]:LIMit:SEGMent[1¦2¦12]:DISTance:STOP <num> | Sets the End Distance for the specified limit segment. | B | T | :CALCulate[1-4][:SELected]:LIMit:DATA | :CALCulate[1-4][:SELected]:LIMit:DATA | Limit
CALCulate[1¦2]:LIMit:SEGMent[1¦2¦12]:FREQuency:STARt <num> | Sets the Begin Frequency for the specified limit segment. | B | T | :CALCulate[1-4][:SELected]:LIMit:DATA | :CALCulate[1-4][:SELected]:LIMit:DATA | Limit
CALCulate[1¦2]:LIMit:SEGMent[1¦2¦12]:FREQuency:STOP <num> | Sets the End Frequency for the specified limit segment. | B | T | :CALCulate[1-4][:SELected]:LIMit:DATA | :CALCulate[1-4][:SELected]:LIMit:DATA | Limit
CALCulate[1¦2]:LIMit:SEGMent[1¦2¦12]:POWer:STARt <num> | Sets the Begin Power for the specified limit segment. | B | T | :CALCulate[1-4][:SELected]:LIMit:DATA | :CALCulate[1-4][:SELected]:LIMit:DATA | Limit
CALCulate[1¦2]:LIMit:SEGMent[1¦2¦12]:POWer:STOP <num> | Sets the End Power for the specified limit segment. | B | T | :CALCulate[1-4][:SELected]:LIMit:DATA | :CALCulate[1-4][:SELected]:LIMit:DATA | Limit
CALCulate[1¦2]:LIMit:SEGMent[1¦2¦12]:STATe <ON¦OFF> | Turns the specified limit segment on/off | B | T | :CALCulate[1-4][:SELected]:LIMit:DATA | :CALCulate[1-4][:SELected]:LIMit:DATA | Limit
CALCulate[1¦2]:LIMit:SEGMent[1¦2¦12]:TYPE <char> | Sets the limit type for the specified segment, Max line/Min line/Max point/Min point. | B | T | :CALCulate[1-4][:SELected]:LIMit:DATA | :CALCulate[1-4][:SELected]:LIMit:DATA | Limit
CALCulate[1¦2]:LIMit:STATe <ON¦OFF> | Turns the limit test on/off. | A+ | T | :CALCulate[1-4][:SELected]:LIMit[:STATe] | :CALCulate[1-4][:SELected]:LIMit[:STATe] | Limit
CALCulate[1¦2]:MARKer:AOFF | Turns off all markers for a given channel | B | T | :CALCulate[1-4][:SELected]:MARKer[1-10][:STATe] | :CALCulate[1-4][:SELected]:MARKer[1-10][:STATe] |  | Need to turn off each marker independently
CALCulate[1¦2]:MARKer:BWIDth <num> | Calculates the bandwidth of a bandpass filter | B | T | :CALCulate[1-４][:SELected]:MARKer[1-10]:BWIDth:DATA? | :CALCulate[1-４][:SELected]:MARKer[1-10]:BWIDth:DATA? |  | Need to turn ON bund search function with ":CALC:[1-4]:MARK:BAND ON", and set band width with ":CALC[1-4]:MARK[1-10]:BWID:THR
CALCulate[1¦2]:MARKer:FUNCtion:RESult? | Queries the results of the active marker function | B | T
CALCulate[1¦2]:MARKer:FUNCtion[:SELect] <char> | Selects the active marker function (Max, Min, Tgt, BW, Notch Multi-peak, Multi-notch, Statistic, Flatness, FilterStatics) | C | T | For MAX,MIN,PEAK,LPEak,RPEak,TARGet,LTARget,RTARget :CALCulate[1-4][:SELected]:MARKer[1-10]:FUNCtion:TYPE For Multi-peak,Multi-notch :CALCulate[1-4][:SELected]:MARKer[1-10]:FUNCtion:MULTi:PPOLarity <POSitive¦NEGative¦BOTH> For Statisitic :CALCulate[1-4]:MST ON, :CALCulate[1-4]:MST:DATA? For Flatness :CALCulate[1-4]:[:SELected]:MARKer:MATH:FLATness[:STATe] ON, :CALCulate[1-4]:[:SELected]:MARKer:MATH:FLATness:DATA? For FilterStatics :CALCulate[1-4]:[:SELected]:MARKer:MATH:FSTatistics[:STATe] ON, :CALCulate[1-4]:[:SELected]:MARKer:MATH:FSTatistics:DATA? | For MAX,MIN,PEAK,LPEak,RPEak,TARGet,LTARget,RTARget :CALCulate[1-4][:SELected]:MARKer[1-10]:FUNCtion:TYPE For Multi-peak,Multi-notch :CALCulate[1-4][:SELected]:MARKer[1-10]:FUNCtion:MULTi:PPOLarity <POSitive¦NEGative¦BOTH> For Statisitic :CALCulate[1-4]:MST ON, :CALCulate[1-4]:MST:DATA? For Flatness :CALCulate[1-4]:[:SELected]:MARKer:MATH:FLATness[:STATe] ON, :CALCulate[1-4]:[:SELected]:MARKer:MATH:FLATness:DATA? For FilterStatics :CALCulate[1-4]:[:SELected]:MARKer:MATH:FSTatistics[:STATe] ON, :CALCulate[1-4]:[:SELected]:MARKer:MATH:FSTatistics:DATA?
CALCulate[1¦2]:MARKer:FUNCtion:TRACking<ON¦OFF> | Turn marker function tracking on/off. | A+ | T | :CALCulate[1-4][:SELected]:MARKer[1-10]:FUNCtion:TRACking | :CALCulate[1-4][:SELected]:MARKer[1-10]:FUNCtion:TRACking
CALCulate[1¦2]:MARKer[1¦2¦8]:GDELay? | Returns the group delay value at the specified marker. | B | T | :CALCulate[1-4][:SELected]:FORMat GDELay :CALC[1-4]:MARK[1-10] ON :CALC[1-4]:MARK[1-10]:Y? | :CALCulate[1-4][:SELected]:FORMat GDELay :CALC[1-4]:MARK[1-10] ON :CALC[1-4]:MARK[1-10]:Y? |  | Need to select Gdelay format and turn a marker ON
CALCulate[1¦2]:MARKer[1¦2¦8]:MAXimum | Sets the specified marker to the maximum value on the trace. | B | T | :CALCulate[1-4][:SELected]:MARKer[1-10]:FUNCtion:EXECute | :CALCulate[1-4][:SELected]:MARKer[1-10]:FUNCtion:EXECute |  | Need to select ":CALCulate[1-16][:SELected]:MARKer[1-10]:FUNCtion:TYPE MAX" first
CALCulate[1¦2]:MARKer[1¦2¦8]:MAXimum:LEFT | Moves the marker to the next local maximum to the left. | B | T | :CALC[1-4]:MARK[1-10]:FUNC:PPOL POS :CALC[1-4]:MARK[1-10]:FUNC:TYPE LPEak :CALC[1-4]:MARK[1-10]:FUNC:EXEC | :CALC[1-4]:MARK[1-10]:FUNC:PPOL POS :CALC[1-4]:MARK[1-10]:FUNC:TYPE LPEak :CALC[1-4]:MARK[1-10]:FUNC:EXEC |  | Need to select peak search type and polarity first
CALCulate[1¦2]:MARKer[1¦2¦8]:MAXimum:RIGHt | Moves the marker to the next local maximum to the right. | B | T | :CALC[1-4]:MARK[1-10]:FUNC:PPOL POS :CALC[1-4]:MARK[1-10]:FUNC:TYPE RPEak :CALC[1-4]:MARK[1-10]:FUNC:EXEC | :CALC[1-4]:MARK[1-10]:FUNC:PPOL POS :CALC[1-4]:MARK[1-10]:FUNC:TYPE RPEak :CALC[1-4]:MARK[1-10]:FUNC:EXEC |  | Need to select peak search type and polarity first
CALCulate[1¦2]:MARKer[1¦2¦8]:MINimum | Sets the marker to the minimum value on the trace. | B | T | :CALCulate[1-4][:SELected]:MARKer[1-10]:FUNCtion:EXECute | :CALCulate[1-4][:SELected]:MARKer[1-10]:FUNCtion:EXECute |  | Need to select ":CALCulate[1-16][:SELected]:MARKer[1-10]:FUNCtion:TYPE MIN" first
CALCulate[1¦2]:MARKer[1¦2¦8]:MINimum:LEFT | Moves the marker to the next local minimum to the left. | B | T | :CALC[1-4]:MARK[1-10]:FUNC:PPOL NEG :CALC[1-4]:MARK[1-10]:FUNC:TYPE LPEak :CALC[1-4]:MARK[1-10]:FUNC:EXEC | :CALC[1-4]:MARK[1-10]:FUNC:PPOL NEG :CALC[1-4]:MARK[1-10]:FUNC:TYPE LPEak :CALC[1-4]:MARK[1-10]:FUNC:EXEC |  | Need to select peak search type and polarity first
CALCulate[1¦2]:MARKer[1¦2¦8]:MINimum:RIGHt | Moves the marker to the next local minimum to the right. | B | T | :CALC[1-4]:MARK[1-10]:FUNC:PPOL NEG :CALC[1-4]:MARK[1-10]:FUNC:TYPE RPEak :CALC[1-4]:MARK[1-10]:FUNC:EXEC | :CALC[1-4]:MARK[1-10]:FUNC:PPOL NEG :CALC[1-4]:MARK[1-10]:FUNC:TYPE RPEak :CALC[1-4]:MARK[1-10]:FUNC:EXEC |  | Need to select peak search type and polarity first
CALCulate[1¦2]:MARKer:MODE <char> | Turns delta marker state on/off, ABSolute/RELative | A | T | :CALCulate[1-4][:SELected]:MARKer[1-10]:REFerence[:STATe] | :CALCulate[1-4][:SELected]:MARKer[1-10]:REFerence[:STATe] |  | ABSolute = OFF RELative = ON
CALCulate[1¦2]:MARKer:NOTCh <num> | Calculates the notch width of a notch filter | B | T | :CALCulate[1-４][:SELected]:MARKer[1-10]:BWIDth:DATA? | :CALCulate[1-４][:SELected]:MARKer[1-10]:BWIDth:DATA? |  | Need to turn ON a band width search feature ON with ":CALC:MARK:BWED ON", and set band width value with "CALC:MARK:BWID:THR" first
CALCulate[1¦2]:MARKer[1¦2¦8]:POINt | Sets the specified marker point. | D | F
CALCulate[1¦2]:MARKer:REFerence:X? | Queries the frequency of the reference marker. | A | T | :CALCulate[1-4][:SELected]:MARKer10:X | :CALCulate[1-4][:SELected]:MARKer10:X |  | Marker 10 = Ref marker
CALCulate[1¦2]:MARKer:REFerence:Y? | Queries the amplitude of the reference marker. | A | T | :CALCulate[1-4][:SELected]:MARKer10:Y | :CALCulate[1-4][:SELected]:MARKer10:Y |  | Marker 10 = Ref marker
CALCulate[1¦2]:MARKer[1¦2¦8][:STATe] <ON¦OFF> | Turns the specified marker on/off. | A+ | T | :CALCulate[1-4][:SELected]:MARKer[1-10][:STATe] | :CALCulate[1-4][:SELected]:MARKer[1-10][:STATe]
CALCulate[1¦2]:MARKer[1¦2¦8]:TARGet <char>,<num> | Performs a marker search for a target value | C | T | :CALCulate[1-4][:SELected]:MARKer[1-10]:FUNCtion:TYPE :CALCulate[1-4][:SELected]:MARKer[1-10]:FUNCtion:EXEC | :CALCulate[1-4][:SELected]:MARKer[1-10]:FUNCtion:TYPE :CALCulate[1-4][:SELected]:MARKer[1-10]:FUNCtion:EXEC |  | Chose "RTARget" for right direction search,"LTARget" for left
CALCulate[1¦2]:MARKer[1¦2¦8]:X <num> | Sets the specified marker frequency (or power in power sweep). | A+ | T | :CALCulate[1-4][:SELected]:MARKer[1-10]:X | :CALCulate[1-4][:SELected]:MARKer[1-10]:X |  | Unit: Hz/dBm/s
CALCulate[1¦2]:MARKer[1¦2¦8]:X:ABS <num> | Sets a marker to an absolute value. | D | F
CALCulate[1¦2]:MARKer[1¦2¦8]:Y? | Queries the specified marker amplitude. | A+ | T | :CALCulate[1-4][:SELected]:MARKer[1-10]:Y? | :CALCulate[1-4][:SELected]:MARKer[1-10]:Y?
CALCulate[1¦2]:MARKer[1¦2¦8]:Y:INDuctance? | Queries the marker's inductance in Smith chart format. | B | T | :CALCulate[1-4][:SELected]:MARKer[1-10]:Y? | :CALCulate[1-4][:SELected]:MARKer[1-10]:Y? |  | Need to calculate inductance value from reactance value of R + jX format result. ":CALC:FORM:SMITh for (R+jX), SADMittance for (G+jB)
CALCulate[1¦2]:MARKer[1¦2¦8]:Y:MAGNitude? | Queries the marker's magnitude in polar format. | B | T | :CALCulate[1-4][:SELected]:MARKer[1-10]:Y? | :CALCulate[1-4][:SELected]:MARKer[1-10]:Y? |  | Need to select "CALC[1-4]:FORM PLINear/PLOGaristmic" first
CALCulate[1¦2]:MARKer[1¦2¦8]:Y:PHASe? | Queries the marker's phase value when in polar format. | B | T | :CALCulate[1-4][:SELected]:MARKer[1-10]:Y? | :CALCulate[1-4][:SELected]:MARKer[1-10]:Y? |  | Need to select ":CALC[1-4]:FORM PLINear/PLOGaristmic" first
CALCulate[1¦2]:MARKer[1¦2¦8]:Y:REACtance? | Queries the marker's reactance value in Smith chart format. | B | T | :CALC[1-4]:FORM SMITh :CALCulate[1-4][:SELected]:MARKer[1-10]:Y? | :CALC[1-4]:FORM SMITh :CALCulate[1-4][:SELected]:MARKer[1-10]:Y? |  | Need to select "CALC[1-4]:FORM SMITh"
CALCulate[1¦2]:MARKer[1¦2¦8]:Y:RESistance? | Queries the marker's resistance value in Smith chart format. | B | T | :CALC[1-4]:FORM SMITh :CALCulate[1-4][:SELected]:MARKer[1-10]:Y? | :CALC[1-4]:FORM SMITh :CALCulate[1-4][:SELected]:MARKer[1-10]:Y? |  | Need to select "CALC[1-4]:FORM SMITh"
CALCulate[1¦2]:MATH[:EXPRession] <expr> | Selects a trace math expression | C | T | :CALCulate[1-4][:SELected]:MATH:FUNCtion | :CALCulate[1-4][:SELected]:MATH:FUNCtion |  | When plural traces are used in eligible channel, need to select an eligible trace by ":CALC[1-4]:PAR[1-4]:SEL" first. Select parameter from "NORM, DIVide, MULTiply, SUBTract, ADD"
CALibration:SELF:ALL | Initiates a SelfCal on all ports. | D | F |  |  |  | Need to repeat ":SENSe[1-4]:MULTiport:PORT[1-2]:SELect"and ":SENSe[1-4]:CORRection:MULTiport:SELF:ONCE"
CALibration:SELF:METHod <ONEPort¦TWOPort> | Selects the method of SelfCal. | D | F |  |  |  | Use the same method defined by ":SENSe[1-4]:CORRection:MULTiport:COLLect:METHod <ERES¦SOLT2>"
CALibration:SELF <ON¦OFF¦ONCE> | Initiates a SelfCal on the currently selected ports. | D | T | :SENSe[1-4]:CORRection:MULTiport:SELF <bool> :SENSe[1-4]:CORRection:MULTiport:SELF:ONCE
CALibration:SELF:TIMER <num> | Sets the time interval for automatic SelfCals. | D | T | :SENSe[1-4]:CORRection:MULTiport:SELF：TIMer
CALibration:ZERO:AUTO <ON¦OFF¦ONCE> | Turns the broadband detector autozeroing function on/off. | D | F |  |  |  | No function
CONFigure <string> | Configures the NA to meas a specific device type & parameter. | D | F |  |  |  | Use control commands or state files
CONTrol[1¦2]:MULTiport:STATE <ON¦OFF> | When on, configures analyzer for use with a multiport test set. | D | T | :SENSe[1-4]:MULTiport[:STATe] <bool>
DIAGnostic:CCONstants:INSTalled? | Queries if correction constants are installed in flash. | D | F |  |  |  | Calibration data is read-only. Status can be checked with ":SENS[1-4]:CORR:COLL:METH:TYPE?" and ":SENS[1-4]:CORR:STAT?"
DIAGnostic:CCONstants:LOAD | Loads default factory cal constants from FD to memory. | D | F |  |  |  | ":SYS:CORR <ON¦OFF>" turns on/off a factory cal data
DIAGnostic:CCONstants:STORE:DISK | Stores default factory cal constants from memory to FD. | D | F |  |  |  | Pre saved in HDD
DIAGnostic:CCONstants:STORE:EEPRom | Stores default factory cal constants from memory to flash. | D | F |  |  |  | Pre saved in HDD
DIAGnostic:COMMunicate:LAN:NETWork:STAT | Displays the menu to set the IP address and to perform ping. | D | F |  |  |  | Only front key access
DIAGnostic:COMMunicate:LAN:PING:IMM | Pings a remote user-specified IP address. | D | F |  |  |  | No equivalent command
DIAGnostic:COMMunicate:LAN:PING:IPAD <string> | STRING Sets the IP address to ping. | D | F |  |  |  | No equivalent command
DIAGnostic:COMMunicate:LAN:SEND <IP_address>,<port_num>,<string>,<timeout> | Instructs NA to open a socket to the IP address & port No. | D | F |  |  |  | Only front key access
DIAGnostic:MDISplay[1¦2]:CORRection <string> | Displays corrected measurement uncertainties. | D | F |  |  |  | No function
DIAGnostic:MDISplay[1¦2]:RESTore | Return to meas mode after viewing cal uncertainties. | D | F |  |  |  | No function
DIAGnostic:PORT:READ? <port><register> 1 | Reads the rear panel I/O ports. | C | T | :CONT:HAND:[C-E] ? | :CONT:HAND:[C-E] ? |  | Need to select input mode by ":CONT:HAND:[C-D]:MODE INPut"
DIAGnostic:PORT:WRITE <port><register> 2 | Writes to the rear panel I/O ports. | C | T | :CONT:HAND[A-F] :CONT:HAND:OUTP[1-2] | :CONT:HAND[A-F] :CONT:HAND:OUTP[1-2] |  | Need to select output mode by ":CONT:HAND:[C-D]:MODE OUTPut"
DIAGnostic:SNUMber <string>? | Queries the instrument's serial number. | B | T | *IDN? | *IDN? |  | Return value is "Agilent Technologies,{Model No. 6char},{Serial No. 10 chara},{FW Rev 5 chara}"
DIAGnostic:SPUR:METHod <NONE¦DITHer¦AVOid> | Selects the spur avoid mode. | D | F |  |  |  | Always ON
DISPlay:ANNotation:CHANnel[1¦2][:STATe] <OFF¦ON> | Enables/disables measurement channel annotation. | C | T | :DISP:WIND[1-4]:TITL | :DISP:WIND[1-4]:TITL |  | Use tｉｔｌｅ Label
DISPlay:ANNotation:CHANnel[1¦2]:USER:LABel:DATA <string> | Specifies the string to be displayed in the annotation area. | C | T | :DISP:WIND[1-4]:TITL:DATA | :DISP:WIND[1-4]:TITL:DATA |  | Use tｉｔｌｅ Label
DISPlay:ANNotation:CHANnel[1¦2]:USER:STATe<OFF¦ON> | Enables user-defined measurement channel annotation. | C | T | :DISP:WIND[1-4]:TITL | :DISP:WIND[1-4]:TITL |  | Use tｉｔｌｅ Label
DISPlay:ANNotation:CLOCk:DATE:FORMat <char> | Selects the Year/Month/Day ordering of the date. | D | F |  |  |  | ":DISP:CLOC" can turn ON/Off the clock display
DISPlay:ANNotation:CLOCk:DATE:MODE <char> | Selects the format for the date in the clock display. | D | F |  |  |  | ":DISP:CLOC" can turn ON/Off the clock display
DISPlay:ANNotation:CLOCk:MODE <char> | Selects how the clock will appear in the meas display title area. | D | F |  |  |  | ":DISP:CLOC" can turn ON/Off the clock display
DISPlay:ANNotation:CLOCk:SEConds[:STATe]<ON¦OFF> | Turns on/off display of seconds in the clock display. | D | F |  |  |  | ":DISP:CLOC" can turn ON/Off the clock display
DISPlay:ANNotation:FREQuency[1¦2]:MODE <char> | Sets the frequency annotation on the display. | D | F
DISPlay:ANNotation:FREQuency[1¦2]:RESolution <char> | Sets the resolution of display frequency values. | D | F
DISPlay:ANNotation:FREQuency [1¦2] [:STATe] <OFF¦ON> | Enables/disables frequency annotation. | A | T | :DISPlay:ANNotation:FREQuency[:STATe] | :DISPlay:ANNotation:FREQuency[:STATe]
DISPlay:ANNotation:FREQuency[1¦2]:USER:LABel:DATA <string> | A user-defined X-axis label. | D | F
DISPlay:ANNotation:FREQuency[1¦2]:USER:STARt<num> | Specifies the start value for user-defined frequency annotation. | D | F
DISPlay:ANNotation:FREQuency[1¦2]:USER:STATe[OFF¦ON] | Enables user-defined frequency annotation. | D | F
DISPlay:ANNotation:FREQuency[1¦2]:USER:STOP<num> | Specifies the stop value for user-defined frequency annotation. | D | F
DISPlay:ANNotation:FREQuency[1¦2]:USER:SUFFix[:DATA] <string> | Specifies the suffix for user defined frequency annotation. | D | F
DISPlay:ANNotation:LIMit:ICON[1¦2]:FLAG[:STATe] <ON¦OFF> | Enables/disables the display of the limit test fail icon. | B | T | :DISP:FSIG {ON¦OFF¦1¦0} | :DISP:FSIG {ON¦OFF¦1¦0} |  | Need to turn ON each channel active trace's limit test by the :CALC:{1-4}:LIM command.
DISPlay:ANNotation:LIMit:ICON[1¦2] :POSition:X <num> | Positions the limit test pass/fail text and icon on the display. | D | F
DISPlay:ANNotation:LIMit:ICON[1¦2] :POSition:Y <num> | Positions the limit test pass/fail text and icon on the display. | D | F
DISPlay:ANNotation:LIMit:ICON[1¦2] :TEXT[:STATe] <ON¦OFF> | Turns the limit test "PASS/FAIL" text on or off. | A | T | :DISP:FSIG | :DISP:FSIG
DISPlay:ANNotation:MARKer[1¦2]:NUMBers [:STATe] <OFF¦ON> | Enables/disables the display of mkr numbers on trace mkrs. | D | F
DISPlay:ANNotation:MARKer[1¦2][:STATe]<ON¦OFF> | Enables/disables the active mkr annotation for meas ch 1 & 2. | D | F
DISPlay:ANNotation:MESSage:AOFF | Turns off any currently showing message window | B | T | :DISP:TABL OFF | :DISP:TABL OFF |  | Use "ECHO Window"
DISPlay:ANNotation:MESSage:CLEar | Removes a user-defined pop-up message window. | B | T | :DISP:ECHO:CLE | :DISP:ECHO:CLE |  | Use "ECHO Window"
DISPlay:ANNotation:MESSage:DATA <string> | Displays a user-defined message. | B | T | :DISP:ECHO | :DISP:ECHO |  | Use "ECHO Window"
DISPlay:ANNotation:MESSage:STATe <ON¦OFF> | Enables/disables the message window | B | T | :DISP:TABL:TYPE ECHO :DISP:TABL | :DISP:TABL:TYPE ECHO :DISP:TABL |  | Use "ECHO Window"
DISPlay:ANNotation:TITLe[1¦2]:DATA <string> | Enters a string for the specified title line. | A | T | :DISPlay:WINDow[1-4]:TITLe:DATA | :DISPlay:WINDow[1-4]:TITLe:DATA
DISPlay:ANNotation:TITLe[:STATe] <ON¦OFF> | Turns on/off display of the title and clock. | B | T | :DISPlay:WINDow[1-4]:TITLe[:STATe] :DISP:CLOC | :DISPlay:WINDow[1-4]:TITLe[:STATe] :DISP:CLOC |  | Clock is controlled independently
DISPlay:ANNotation:YAXis:MODE <char> | Sets mode for the Y-axis labels | D | F
DISPlay:ANNotation:YAXis[:STATe] <ON¦OFF> | Turns on/off Y-axis labels. | A | T | :DISPlay:WINDow[1-4]:LABel | :DISPlay:WINDow[1-4]:LABel
DISPlay:CMAP:COLor[1¦2¦・6]:GREYscale <num> | Changes the default intensity of the selected item. | B | T | :DISP:COL[1-2]:BACK :DISP:COL[1-2]:GRAT[1-2] :DISP:COL[1-2]:LIM[1-2] :DISP:COL[1-2]:TRAC[1-4]:DATA :DISP:COL[1-2]:TRAC[1-4]:MEM | :DISP:COL[1-2]:BACK :DISP:COL[1-2]:GRAT[1-2] :DISP:COL[1-2]:LIM[1-2] :DISP:COL[1-2]:TRAC[1-4]:DATA :DISP:COL[1-2]:TRAC[1-4]:MEM |  | Use RGB format
DISPlay:CMAP:COLor[1¦2¦・6]:RGB <num,num,num> | Sets the color map based on the R/G/B model. (Ext monitor) | B | T | :DISP:COL[1-2]:BACK :DISP:COL[1-2]:GRAT[1-2] :DISP:COL[1-2]:LIM[1-2] :DISP:COL[1-2]:TRAC[1-4]:DATA :DISP:COL[1-2]:TRAC[1-4]:MEM | :DISP:COL[1-2]:BACK :DISP:COL[1-2]:GRAT[1-2] :DISP:COL[1-2]:LIM[1-2] :DISP:COL[1-2]:TRAC[1-4]:DATA :DISP:COL[1-2]:TRAC[1-4]:MEM |  | Use RGB format
DISPlay:CMAP:DEFault | Sets the color scheme to the factory default. (Ext monitor) | A | T | :DISPlay:COLor[1-2]:RESet | :DISPlay:COLor[1-2]:RESet
DISPlay:CMAP:SCHeme | Sets the color scheme for an external monitor. | B | T | :DISP:COL[1-2]:BACK :DISP:COL[1-2]:GRAT[1-2] :DISP:COL[1-2]:LIM[1-2] :DISP:COL[1-2]:TRAC[1-4]:DATA :DISP:COL[1-2]:TRAC[1-4]:MEM | :DISP:COL[1-2]:BACK :DISP:COL[1-2]:GRAT[1-2] :DISP:COL[1-2]:LIM[1-2] :DISP:COL[1-2]:TRAC[1-4]:DATA :DISP:COL[1-2]:TRAC[1-4]:MEM |  | Use RGB format
DISPlay:FORMat <char> | Selects the format (full or split) for displaying trace data | B | T | :DISPlay:SPLit | :DISPlay:SPLit |  | Available only in the same channel
DISPlay:FORMat:EXPAND <ON¦OFF> | Enables/disables expand measurement mode. | A | T | :DSP:SKEY OFF | :DSP:SKEY OFF
DISPlay:MENU:KEY[1¦2¦7] <string> | Specifies the softkey labels from a controller or IBASIC | D | F |  |  |  | Use VBA
DISPlay:MENU[2]:KEY[1¦2¦7] <string> | Specifies the softkey labels when using user-defined BEGIN key. | D | F |  |  |  | Use VBA
DISPlay:MENU:RECall:FAST[:STATe] <ON¦OFF> | Turns on/off fast recall mode. | D | F |  |  |  | Always ON
DISPlay:PROGram[:MODE] <char> | Selects the portion of the screen to be used as IBASIC | D | F |  |  |  | Cannot control VBA editor display
DISPlay:WINDow[1¦2¦10]:GEOMetry:LLEFT? | Queries the absolute pixel coordinates of the lower left corner. | D | F
DISPlay:WINDow[1¦2¦10]:GEOMetry:SIZE? | Queries the width and height (in pixels). | D | F
DISPlay:WINDow[1¦2¦10]:GEOMetry:URIGHT? | Queries the absolute pixel coordinates of the upper right corner. | D | F
DISPlay:WINDow:GRAPhics:BUFFer[:STATe]<ON¦OFF> | Turn on/off buffering of user graphics commands. | D | F
DISPlay:WINDow[1¦2¦10]:GRAPhics :CIRCle <num> | Draws a circle of the specified Y-axis radius | D | F
DISPlay:WINDow[1¦2¦10]:GRAPhics :CLEar | Clears the user graphics and graphics buffer. | D | F
DISPlay:WINDow[1¦2¦10]:GRAPhics :COLor <num> | Sets the color of the user graphics pen | D | F
DISPlay:WINDow[1¦2¦10]:GRAPhics [:DRAW] <num1>,<num2>3 | Draws a line | D | F
DISPlay:WINDow[1¦2¦10]:GRAPhics:LABel <string> | Draws a label with the lower left corner. | D | F
DISPlay:WINDow[1¦2¦10]:GRAPhics:LABel:FONT <char> | Selects the user graphics label font | D | F
DISPlay:WINDow[1¦2¦10]:GRAPhics :MOVE <num1>,<num2> | Moves the pen to the specified new pen position | D | F
DISPlay:WINDow[1¦2¦10]:GRAPhics:RECTangle <num1>,<num2> | Draws a rectangle of the specified size with lower left corner | D | F
DISPlay:WINDow[1¦2¦10]:GRAPhics:SCALe <xmin>,<xmax>,<ymin,<ymax> | Specifies new coordinates for window. | D | F
DISPlay:WINDow[1¦2¦10]:GRAPhics :STATe? | Queries whether a window is enabled for user graphics commands. | D | F
DISPlay:WINDow[1¦2]:TRACe:GRATicule:GRID[:STATe]<ON¦OFF> | Turns display graticule on/off. | D | F
DISPlay:WINDow[1¦2]:TRACe[1¦2][:STATe]<ON¦OFF> | Turns the display of trace and memory data on/off. | B | T | :DISPlay:WINDow[1-4]:TRACe[1-4]:STATe :DISPlay:WINDow[1-4]:TRACe[1-4]:MEM | :DISPlay:WINDow[1-4]:TRACe[1-4]:STATe :DISPlay:WINDow[1-4]:TRACe[1-4]:MEM |  | Need to specify DATA and MEMory
DISPlay:WINDow[1¦2]:TRACe:Y[:SCALe]:AUTO ONCE | Scales the measurement data for a best fit display. | A | T | :DISPlay:WINDow[1-4]:TRACe[1-4]:Y[:SCALe]:AUTO | :DISPlay:WINDow[1-4]:TRACe[1-4]:Y[:SCALe]:AUTO
DISPlay:WINDow[1¦2]:TRACe:Y[:SCALe]:PDIVision<num> | Specifies the height of each vertical division. | A+ | T | :DISPlay:WINDow[1-4}:TRACe[1-4}:Y[:SCALe]:PDIVision | :DISPlay:WINDow[1-4}:TRACe[1-4}:Y[:SCALe]:PDIVision
DISPlay:WINDow[1¦2]:TRACe:Y[:SCALe]:RLEVel<num> | Specifies the value for the Y-axis reference position | A+ | T | :DISPlay:WINDow[1-4}:TRACe[1-4}:Y[:SCALe]:RLEVel | :DISPlay:WINDow[1-4}:TRACe[1-4}:Y[:SCALe]:RLEVel
DISPlay:WINDow[1¦2]:TRACe:Y[:SCALe]:RPOSition <num> | Specifies the Y-axis reference position | A+ | T | :DISPlay:WINDow[1-4}:TRACe[1-4}:Y[:SCALe]:RPOSition | :DISPlay:WINDow[1-4}:TRACe[1-4}:Y[:SCALe]:RPOSition
DISPlay:WINDow[1¦2¦10]:TRACe[1¦2]:Y:TRACk <OFF¦PEAK¦FREQ> | Selects the method for reference offset tracking. | D | F
DISPlay:WINDow[1¦2¦10]:TRACe[1¦2]:Y:TRACk:FREQuency <num> | Selects frequency to track with reference tracking. | D | F
FORMat:BORDer <char> | Specifiesa the byte order used for HPIB data transfer | A+ | T | :FORMat:BORDer | :FORMat:BORDer | Data
FORMat[:DATA] <char>[,<num>] | Specifies the data format for use during data transfer | C | T | :FORMat:DATA | :FORMat:DATA | Data | "Integer" is not available
HCOPy:ABORt | Aborts any hardcopy currently in progress. | A+ | T | :HCOPy:ABORt | :HCOPy:ABORt
HCOPy:DEVice[1¦2¦3]:COLor <ON¦OFF> | Selects color ON / OFF for hardcopy output. | D | F
HCOPy:DEVice[1¦2¦3]:LANGuage <char> | Selects language for hardcopy output. | D | F
HCOPy:DEVice[1¦2¦3]:MODE <char> | Selects the graph and/or table(s) to appear on a hardcopy plot | D | F
HCOPy:DEVice[1¦2¦3]:PAGE:MARGin:LEFT <num> | Sets the left margin (for printer output) in millimeters. | D | F
HCOPy:DEVice[1¦2¦3]:PAGE:MARGin:TOP <num> | Sets the top margin (for printer output) in millimeters. | D | F
HCOPy:DEVice[1¦2¦3]:PAGE:ORIentation <char> | Sets printer output page orientation | D | F
HCOPy:DEVice[1¦2¦3]:PAGE:WIDTh <num> | Sets the print width (for printer output) in millimeters. | D | F
HCOPy:DEVice[1¦2¦3]:PORT <char> | Selects the communications port for hardcopy output | D | F
HCOPy:DEVice[1¦2]:RESolution <num> | Sets the printer resolution in dots per inch. | D | F
HCOPy[:IMMediate] | Initiates a hardcopy output (print or plot). | A+ | T | :HCOPy[:IMMediate] | :HCOPy[:IMMediate]
HCOPy:ITEM[1¦2¦3]:ANNotation:STATe <ON¦OFF> | Turns on/off ch & freq annotation as part of hardcopy output. | D | F
HCOPy:ITEM[1¦2¦3]:FFEed:STATe <ON¦OFF> | Turns on/off an automatic form feed | D | F
HCOPy:ITEM[1¦2¦3]:GRATicule:STATe <ON¦OFF> | Turns on/off graticule as part of hardcopy output. | D | F
HCOPy:ITEM[1¦2¦3]:MARKer:STATe <ON¦OFF> | Turns on/off marker symbols as part of hardcopy output. | D | F
HCOPy:ITEM[1¦2¦3]:TITLe:STATe <ON¦OFF> | Turns on/off title and clock lines as part of hardcopy output. | D | F
HCOPy:ITEM[1¦2¦3]:TRACe:STATe <ON¦OFF> | Turns on/off trace data as part of hardcopy output. | D | F
INITiate[1¦2]:CONTinuous <ON¦OFF> | Sets the trigger system to continuous sweep or to stop sweep. | A+ | T | :INITiate[1-4]:CONTinuous | :INITiate[1-4]:CONTinuous | Trigger
INITiate[1¦2][:IMMediate] | Initiates a new measurement sweep. | C | T | :INITiate[1-4][:IMMediate] :TRIG:SING | :INITiate[1-4][:IMMediate] :TRIG:SING | Trigger | In ENA-L, ":TRIG:SING" is necessary to detect sweep end
INPut:GAIN:AUTo <R¦A¦B>,<ON¦OFF> | Sets the R, A, or B IF input automatic gain control on or off. | D | F
INPut:GAIN:SETTing <R¦A¦B>,<HIGH¦MEDHIGH¦MEDIUM¦LOW> | Sets the R, A, or B IF input gain | D | F
MMEMory:CATalog?<string> | Lists the names of the files in memory. | A+ | T | :MMEMory:CATalog? | :MMEMory:CATalog? |  | Specify a directory name
MMEMory:CDIRectory <string> | Changes the current directory on a DOS formatted disk | D | F |  |  |  | Directly assign with ":MMEMory:CATalog?"
MMEMory:COPY <string1>,<string2> | Copies a file | A+ | T | :MMEMory:COPY | :MMEMory:COPY
MMEMory:DELete <string> | Deletes a file | A+ | T | :MMEMory:DELete | :MMEMory:DELete
MMEMory:INITialize[<string>[,<char>[,<num>]]] | Formats a disk | D | F | :MMEMory:LOAD[:STATe] | :MMEMory:LOAD[:STATe]
MMEMory:LOAD:STATe 1, <string> | Recalls an instrument state from mass storage | A+ | T | :MMEMory:LOAD[:STATe] | :MMEMory:LOAD[:STATe]
MMEMory:LOAD:CKIT:USER[:SELect] <KIT1¦KIT2...KIT10> | Recalls a user cal kit. | B | T | :SENS[1-4]:CORR:COLL:CKIT | :SENS[1-4]:CORR:COLL:CKIT |  | Select from predefined user kit 1 to 10
MMEMory:FILE:INFO?<string> | Returns file information such as date/time. | D | F
MMEMory:MDIRectory <string> | Makes a new directory on a DOS formatted disk. | A+ | T | :MMEM:MDIR | :MMEM:MDIR |  | A:\: Internal Floppy disk drive D:\: Internal hard disk drive
MMEMory:MOVE <string1>,<string2> | Moves or renames a file | B | T | :MMEM:COPY :MMEM:DEL | :MMEM:COPY :MMEM:DEL |  | Need to use copy and delete function
MMEMory:MSIS <string> | Selects a mass storage device | D | F |  |  |  | need to specify in each command
MMEMory:RDIRectory <string> | Deletes a directory from a DOS formatted disk. | A | T | ：MMEM:DEL | ：MMEM:DEL
MMEMory:STORe:STATe 1,<string> | Saves an instrument state to mass storage | A | T | :MMEMory:STORe[:STATe] | :MMEMory:STORe[:STATe] |  | The file named "autorec.sta" is used every power up time
MMEMory:STORe:STATe:CORRection <ON¦OFF> | Turns the calibration on/off | C | T | :MMEMory:STORe:STYPe | :MMEMory:STORe:STYPe |  | Select "CSTate"or "CDSTate" parameters for ON "STATe" or ]"DSTate" for OFF
MMEMory:STORe:STATe:FORMat <char> | Saves instrument state files | D | F |  |  |  | it is impossible to save with the older model's format
MMEMory:STORe:STATe:ISTate <ON¦OFF> | Turns the instrument state on/off | D | F |  |  |  | not available
MMEMory:STORe:STATe:TRACe <ON¦OFF> | Turns the data trace on/off | C | T | :MMEMory:STORe:STYPe | :MMEMory:STORe:STYPe |  | Select "DSTate' or "CDSTate" for ON, "STATe" or "CSTate" for OFF
MMEMory:STORe:STATe:TSCAL <ON¦OFF> | When on, the saved state will be the test set cal only. | D | F |  |  |  | not available. (ENA-L has enough speed and disk volume)
MMEMory:STORe:TRACe <char>,<string> | Stores an ASCII list of trace and frequency values to a file | B | T | :MMEM:STOR::FDAT | :MMEM:STOR::FDAT |  | Need to select a target channel and trace first with ":DISP:WIND[1-4]:ACT" and ":CALC[1-4]:PAR[1-4]:SEL" commands
MMEMory:STORe:TRACe:FORMat <char> | Selects the format that the ASCII data will be saved in. | D | F |  |  |  | CSV format only. VBA macro can provide Touchstone format.
MEMory:TRANsfer:BDAT <string> [,<block>] | Copies a file to or from the analyzer's disk drive. | A | T | :MMEMory:TRANsfer | :MMEMory:TRANsfer
MMEMory:TRANsfer[:HFS] <string> [,<block>] | Copies a file to or from the analyzer's disk drive. | A | T | :MMEMory:TRANsfer | :MMEMory:TRANsfer |  | Use Hierarchical File System as same as normal PC
OUTPut[:STATe] <ON¦OFF> | Turns RF power from the source on/off. | A+ | T | :OUTPut[:STATe] | :OUTPut[:STATe]
POWer[1¦2]:MODE <char> | Specifies either frequency sweep or power sweep. | B | T | :SENSe[1-4]:SWEep:TYPE POWer | :SENSe[1-4]:SWEep:TYPE POWer
PROGram:CATalog? | Lists the names of the defined IBASIC programs | A+ | T | :PROGram:CATalog? | :PROGram:CATalog?
PROGram[:SELected]:DEFine | BLOCK Downloads an IBASIC program from an controller. | D | F |  |  |  | Use VBA instead of IBASIC
PROGram[:SELected]:DELete:ALL | Deletes all IBASIC programs from the program buffer | D | F |  |  |  | Use VBA instead of IBASIC
PROGram[:SELected]:DELete[:SELected] | Deletes the active IBASIC program | D | F |  |  |  | Use VBA instead of IBASIC
PROGram[:SELected]:EXECute <string> | Executes an IBASIC command. | D | F |  |  |  | Use VBA instead of IBASIC
PROGram[:SELected]:MALLocate <num> | Allocates memory space for IBASIC programs | D | F |  |  |  | Use VBA instead of IBASIC
PROGram[:SELected]:NAME 'PROG' | Selects the IBASIC program in the program buffer to be active. | A+ | T | :PROGram[:SELected]:NAME | :PROGram[:SELected]:NAME |  | Use VBA instead of IBASIC
PROGram[:SELected]:NUMBer <string>,<data> | Loads a value for a numeric variable string in the IBASIC prog | D | F
PROGram[:SELected]:STATe <char> | Selects the state of the active IBASIC program | C | T | :PROGram[:SELected]:STATe | :PROGram[:SELected]:STATe |  | Only "RUN' and "STOP" are available
PROGram[:SELected]:STRing <string1>,<string2> | Loads a value for a string variable string in the IBASIC program | D | F
PROGram[:SELected]:WAIT | Waits until the IBASIC program completes. | D | F
ROUTe[1¦2]:REFLection:PATH:DEFine:PORT<1¦2¦・12> | Selects the test port connected to the REFLECTION port of NA. Multi-port test set only. | D | T | :SENSe[1-4]:MULTiport:PORT1:SEL (port no)
ROUTe[1¦2]:TRANsmission:PATH:DEFine: PORT <1¦2¦・12> | Selects the test port connected to the TRANS port of NA. Multi-port test set only. | D | T | :SENSe[1-4]:MULTiport:PORT2:SEL (port no)
ROUTe[1¦2]:PATH:DEFine: PORT <num>,<num> | Selects the test port for reflection port and transmission port. ES model only | D | T | :SENSe[1-4]:MULTiport:PORT[1-2]:SEL (port no)
SENSe[1¦2]:ACTIVE? | Returns true (1) if the ch is active. | B | T | :SERV:CHAN:ACT?で確認、または :DISP:WIND[1-4]:ACT で指定する | :SERV:CHAN:ACT?で確認、または :DISP:WIND[1-4]:ACT で指定する
SENSe[1¦2]:AVERage:CLEar | Re-starts the trace averaging function. | A+ | T | :SENSe[1-4]:AVERage:CLEar | :SENSe[1-4]:AVERage:CLEar
SENSe[1¦2]:AVERage:COUNt <num> | Specifies a count or weighting factor for the averaged meas data. | A+ | T | :SENSe[1-4]:AVERage:COUNt | :SENSe[1-4]:AVERage:COUNt
SENSe[1¦2]:AVERage[:STATe] <ON¦OFF> | Turns the trace averaging function on/off. | A+ | T | :SENSe[1-4]:AVERage[:STATe] | :SENSe[1-4]:AVERage[:STATe]
SENSe[1¦2]:BWIDth[:RESolution] <num> HZ | Specifies the bandwidth of the IF receiver | A | T | :SENSe[1-4]:BANDwidth[:RESolution] | :SENSe[1-4]:BANDwidth[:RESolution] |  | 871X = 15 (fine) 250 (narrow) 1200 (medium narrow) 3700 (medium) 4000 (medium wide) 6500 (wide) E506XA = 10 Hz to 30 kHz (1,3 step)
SENSe[1¦2]:CORRection:[OFF/ON] | Turns measurement calibration function on/off. | A | T | :SENSe[1-4]:CORRection:STATe | :SENSe[1-4]:CORRection:STATe
SENSe[1¦2]:CORRection:ANNotation? | Returns the current calibration annotation | D | F
SENSe[1¦2]:CORRection:CAPacitance:CONNector <num> | Selects a connector compensating capacitance value. | A | T | :CALCulate[1-4]:SRL:CONNector[1-2]:CAPacitance | :CALCulate[1-4]:SRL:CONNector[1-2]:CAPacitance |  | value should be defined as farad (F).
SENSe[1¦2]:CORRection:CKIT:MODify[:SElect] <char> | Selects or queries the cal kit connector type. (TYPenf¦TYPenm¦UD1¦TYPe35mm¦TYPeff¦UD2¦TYPeapc7) | D | F |  |  |  | Can read calibration kit name by ":SENS[1-4]:CORR:COLL:CKIT:LAB?"
SENSe[1¦2]:CORRection:CKIT:LOAD:MODify:DELay? #0~1e-06[S]# | Sets or queries the cal kit load standard offset delay time. | B | T | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:DEL (/s) | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:DEL (/s) |  | Need to select a calibration kit and LOAD standard by ":SENS[1-4]:CORR:COLL:CKIT" and ":SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:TYPE LOAD" first
SENSe[1¦2]:CORRection:CKIT:LOAD:MODify:LOSS? #0~1e+12# | Sets or queries the cal kit load standard offset loss value. | B | T | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:LOSS (ohm/s) | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:LOSS (ohm/s) |  | Need to select a calibration kit and LOAD standard by ":SENS[1-4]:CORR:COLL:CKIT" and ":SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:TYPE LOAD" first
SENSe[1¦2]:CORRection:CKIT:LOAD:MODify:ZOFFset? #25~100[OHM]# | Sets or queries the cal kit load standard offset impedance value. | B | T | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:Z0 (ohm) | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:Z0 (ohm) |  | Need to select a calibration kit and LOAD standard by ":SENS[1-4]:CORR:COLL:CKIT" and ":SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:TYPE LOAD" first
SENSe[1¦2]:CORRection:CKIT:NAME<KIT1¦KIT2¦...¦KIT10>,<STRING> | Sets the description of the user-defined calibration kit. | B | T | :SENS[1-4]:CORR:COLL:CKIT <1-10> :SENS[1-4]:CORR:COLL:CKIT:LAB | :SENS[1-4]:CORR:COLL:CKIT <1-10> :SENS[1-4]:CORR:COLL:CKIT:LAB |  | Need to select calibration kit first
SENSe[1¦2]:CORRection:CKIT:OPEN:MODify:CONE? #-1e-60~1[FARAD]# | Sets or queries the cal kit open std fringe capacitance C1. | B | T | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:C1 | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:C1 |  | Need to select a calibration kit and OPEN standard by ":SENS[1-4]:CORR:COLL:CKIT" and ":SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:TYPE OPEN" first
SENSe[1¦2]:CORRection:CKIT:OPEN:MODify:CTHRee? #-1e-60~1[FARAD]# | Sets or queries the cal kit open std fringe capacitance C3. | B | T | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:C3 | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:C3 |  | Need to select a calibration kit and OPEN standard by ":SENS[1-4]:CORR:COLL:CKIT" and ":SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:TYPE OPEN" first
SENSe[1¦2]:CORRection:CKIT:OPEN:MODify:CTWO? #-1e-60~1[FARAD]# | Sets or queries the cal kit open std fringe capacitance C2. | B | T | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:C2 | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:C2 |  | Need to select a calibration kit and OPEN standard by ":SENS[1-4]:CORR:COLL:CKIT" and ":SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:TYPE OPEN" first
SENSe[1¦2]:CORRection:CKIT:OPEN:MODify:CZERo? #-1e-60~1[FARAD]# | Sets or queries the cal kit open std fringe capacitance C0. | B | T | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:C0 | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:C0 |  | Need to select a calibration kit and OPEN standard by ":SENS[1-4]:CORR:COLL:CKIT" and ":SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:TYPE OPEN" first
SENSe[1¦2]:CORRection:CKIT:OPEN:MODify:DELay? #0~1e-06[S]# | Modifies or queries the cal kit open std offset delay time. | B | T | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:DEL (/s) | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:DEL (/s) |  | Need to select a calibration kit and OPEN standard by ":SENS[1-4]:CORR:COLL:CKIT" and ":SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:TYPE OPEN" first
SENSe[1¦2]:CORRection:CKIT:OPEN:MODify:LOSS? #0~1e+12# | Modifies or queries the cal kit open std offset loss value. | B | T | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:LOSS (ohm/s) | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:LOSS (ohm/s) |  | Need to select a calibration kit and OPEN standard by ":SENS[1-4]:CORR:COLL:CKIT" and ":SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:TYPE OPEN" first
SENSe[1¦2]:CORRection:CKIT:OPEN:MODify:ZOFFset? #25~100[OHM]# | Modifies or queries the cal kit open std offset impedance value. | B | T | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:Z0 (ohm) | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:Z0 (ohm) |  | Need to select a calibration kit and OPEN standard by ":SENS[1-4]:CORR:COLL:CKIT" and ":SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:TYPE OPEN" first
SENSe[1¦2]:CORRection:CKIT:PRESet[:IMMediate] | Sets all values of all user-defined cal kits to the default values. | B | T | :SENS[1-4]:CORR:COLL:CKIT:RES | :SENS[1-4]:CORR:COLL:CKIT:RES |  | Eligible for selected calibration kit in each channel
SENSe[1¦2]:CORRection:CKIT:SAVE? <KIT1¦KIT2¦...¦KIT10> | Saves the user-defined cal kit. | D | F |  |  |  | Save is done simultaneously with change
SENSe[1¦2]:CORRection:CKIT:SHORt:MODify:DELay? #0~1e-06[S]# | Modifies or queries the cal kit short std offset delay time. | B | T | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:DEL (/s) | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:DEL (/s) |  | Need to select a calibration kit and SHORT standard by ":SENS[1-4]:CORR:COLL:CKIT" and ":SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:TYPE SHORt" first
SENSe[1¦2]:CORRection:CKIT:SHORt:MODify:LOSS? #0~1e+12# | Modifies or queries the cal kit short std offset loss value. | B | T | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:LOSS (ohm/s) | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:LOSS (ohm/s) |  | Need to select a calibration kit and SHORT standard by ":SENS[1-4]:CORR:COLL:CKIT" and ":SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:TYPE SHORt" first
SENSe[1¦2]:CORRection:CKIT:SHORt:MODify:ZOFFset? #25~100[OHM]# | Modifies or queries the cal kit short std offset impedance value. | B | T | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:Z0 (ohm) | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:Z0 (ohm) |  | Need to select a calibration kit and SHORT standard by ":SENS[1-4]:CORR:COLL:CKIT" and ":SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:TYPE SHORt" first
SENSe[1¦2]:CORRection:CKIT:THRU:MODify:DELay? #0~1e-06[S]# | Modifies or queries the cal kit thru std offset delay time. | B | T | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:DEL (/s) | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:DEL (/s) |  | Need to select a calibration kit and THRU standard by ":SENS[1-4]:CORR:COLL:CKIT" and ":SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:TYPE THRU" first
SENSe[1¦2]:CORRection:CKIT:THRU:MODify:LOSS? #0~1e+12# | Modifies or queries the cal kit thru std offset loss value. | B | T | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:LOSS (ohm/s) | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:LOSS (ohm/s) |  | Need to select a calibration kit and THRU standard by ":SENS[1-4]:CORR:COLL:CKIT" and ":SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:TYPE THRU" first
SENSe[1¦2]:CORRection:CKIT:THRU:MODify:ZOFFset? #25~100[OHM]# | Modifies or queries the cal kit thru std offset impedance value. | B | T | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:Z0 (ohm) | :SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:Z0 (ohm) |  | Need to select a calibration kit and THRU standard by ":SENS[1-4]:CORR:COLL:CKIT" and ":SENS[1-4]:CORR:COLL:CKIT:STAN[1-21]:TYPE THRU" first
SENSe[1¦2]:CORRection:CLASs[:SELect]? | Returns the current calibration choice. | D | F |  |  |  | USER CAL only
SENSe[1¦2]:CORRection:CLASs[:SELect] <item> | Selects an existing calibration from Default factory, User cal, etc. | D | F |  |  |  | USER CAL only
SENSe[1¦2]:CORRection:COLLect:ABORt | Aborts the calibration that is currently in progress. | D | F |  |  |  | Need to measure the same pass again, or complete whole process once.
SENSe[1¦2]:CORRection:COLLect[:ACQuire] <char> | Measures a calibration standard (STAN1 to STAN7) | C | T | :SENS[1-4]:CORR:COLL:[ISOL, LOAD, OPEN, SHOR, THRU] | :SENS[1-4]:CORR:COLL:[ISOL, LOAD, OPEN, SHOR, THRU] |  | Need to define necessary standards with ":SENS[1-4]:CORR:COLL:CKIT:ORD:[LOAD,OPEN,SHOR,THRU]"
SENSe[1¦2]:CORRection:COLLect:CKIT:PORT[1¦2¦・12][:SELect] <string> | Selects which port of NA or multiport test set will be assigned a Cal Kit. | D | T | :SENSe[chan]:MULTiport:PORT[1-2]:SELect <port no.> :SENSe[1-4]:CORR:COLL:CKIT |  |  | Need to select channel, calkit, and connected port. Calkit should be defined before for each port type.
SENSe[1¦2]:CORRection:COLLect:CKIT[:SELect] <string> | Selects a Cal Kit | A | T | :SENS[1-4]:CORR:COLL:CKIT | :SENS[1-4]:CORR:COLL:CKIT |  | Use number from 1 to 10 to select a kit
SENSe[1¦2]:CORRection:COLLect:ISTate[:AUTO]<ON¦OFF> | Selects the instrument state for cal, Full Band (ON) or User Defined (OFF). | D | F
SENSe[1¦2]:CORRection:COLLect:METHod <char> | Selects the type of calibration | B | T | :SENS[1-4]:CORR:COLL:METH:{OPEN/SHOR/THRU/ERES/SOLT1/SOLT2} <port>,[<port>] | :SENS[1-4]:CORR:COLL:METH:{OPEN/SHOR/THRU/ERES/SOLT1/SOLT2} <port>,[<port>] |  | OPEN/SHOR/THRU means response calibration with each standard type
SENSe[1¦2]:CORRection:COLLect:MP:OPEN <STAN1¦STAN2¦・STAN12> | Measures an open on the port selected during a test set cal. | C | T | :SENS[1-4]:CORR:MULT:COLL[:ACQuire]:OPEN |  | Multi cal | Specify port number
SENSe[1¦2]:CORRection:COLLect:MP:SHORT <STAN1¦STAN2¦・STAN12> | Measures a short on the port selected during a test set cal. | C | T | :SENS[1-4]:CORR:MULT:COLL[:ACQuire]:SHORt |  | Multi cal | Specify port number
SENSe[1¦2]:CORRection:COLLect:MP:LOAD <STAN1¦STAN2¦・STAN12> | Measures a load on the port selected during a test set cal. | C | T | :SENS[1-4]:CORR:MULT:COLL[:ACQuire]:LOAD |  | Multi cal | Specify port number
SENSe[1¦2]:CORRection:COLLect:MP:THRU <STAN1¦STAN2¦・STAN6> | Measures a thru on the port selected during a test set cal. | C | T | :SENS[1-4]:CORR:MULT:COLL[:ACQuire]:THRU |  | Multi cal | Specify port number
SENSe[1¦2]:CORRection:COLLect:PORTS <2¦4¦6¦8¦10¦12> | Selects the number of ports to perform a test set cal on. | D | F |  |  |  | No pre-define is needed
SENSe[1¦2]:CORRection:COLLect:SAVE | Completes and saves current calibration. | A+ | T | :SENS[1-4]:CORR:COLL:SAVE | :SENS[1-4]:CORR:COLL:SAVE
SENSe[1¦2]:CORRection:COLLect:VERify:REFLection <STAN1¦STAN2¦・STAN12> | Measures a cal std during a cal check procedure for reflection meas. | D | F
SENSe[1¦2]:CORRection:COLLect:VERify:TRANsmission <STAN1¦STAN2¦・STAN12>; | Measures a cal std during a cal check procedure for transmission meas. | D | F
SENSe[1¦2]:CORRection:CSET[:SELect] DEFault | Restores factory default cal for the current meas and ch. | D | F
SENSe[1¦2]:CORRection:CSET[:SELect]? | Queries the current cal type, factory default, full band or user defined. | D | F
SENSe[1¦2]:CORRection:EDELay:TIME <num> | Specifies the electrical delay in seconds. | C | T | :CALC[1-4]:CORR:EDEL:TIME | :CALC[1-4]:CORR:EDEL:TIME |  | Eligible for active trace only
SENSe[1¦2]:CORRection:EXTension:REFLection[:TIME]<num> | Specifies the port extension at the reflection port, in seconds. | C | T | :SENS[1-4]:CORR:EXT:PORT[1-2] | :SENS[1-4]:CORR:EXT:PORT[1-2] |  | Specify port number
SENSe[1¦2]:CORRection:EXTension[:STATe]<ON¦OFF> | Enables port extensions. | A+ | T | :SENS[1-4]:CORR:EXT | :SENS[1-4]:CORR:EXT
SENSe[1¦2]:CORRection:EXTension:TRANsmission[:TIME] <num> | Specifies the port extension at the transmission port. | C | T | :SENS[1-4]:CORR:EXT:PORT[1-2] | :SENS[1-4]:CORR:EXT:PORT[1-2] |  | Specify port number
SENSe[1¦2]:CORRection:IMPedance:INPut:MAGNitude<num> | Specifies the reference impedance for the Smith chart display. | A+ | T | :SENSe[1-4]:CORR:IMP | :SENSe[1-4]:CORR:IMP
SENSe[1¦2]:CORRection:IMPedance:INPut:MAGNitude:SELect ZO_50 | Selects 50 ohms as the system impedance. | A | T | :SENSe[1-4]:CORR:IMP | :SENSe[1-4]:CORR:IMP |  | Use 50 as parameter
SENSe[1¦2]:CORRection:IMPedance:INPut:MAGNitude:SELect ZO_75 | Selects 75 ohms as the system impedance. | A | T | :SENSe[1-4]:CORR:IMP | :SENSe[1-4]:CORR:IMP |  | Use 75 as parameter
SENSe[1¦2]:CORRection:ISOLation[:STATe] {OFF¦0¦ON¦1} | Enables or disables isolation error correction in 2-port cal. | D | F |  |  |  | ISOLATION is always optional
SENSe[1¦2]:CORRection:LENGth:COAX <num> | Specifies the length of cable to be calibrated, in feet or meters. | B | T | :CALC{1-4}:SRL:CONN{1-2}:LENG <num> | :CALC{1-4}:SRL:CONN{1-2}:LENG <num> |  | Use the same command in connector and cable
SENSe[1¦2]:CORRection:LENGth:CONNector <num> | Specifies the length of an interface connector, in mm or inches. | A | T | :CALC[1-4]:SRL:CONN[1-2]:LENG | :CALC[1-4]:SRL:CONN[1-2]:LENG |  | set cable length in meter only
SENSe[1¦2]:CORRection:LOSS:COAX <num> | Specifies the loss of a cable under test, in dB/100 ft. | B | T | :CALC{1-4}:TRAN:DIST:CLOS <num> | :CALC{1-4}:TRAN:DIST:CLOS <num> |  | Active only "distance" is selected with :CALC{1-4}:TRAN:METH DISTcommand. Need to select length unit with ":CALC{1-4}:TRAN:DIST:UNIT {FEET¦MET}"
SENSe[1¦2]:CORRection:MODel:CONNector[:IMMediate] | Measures the cable connector and determine optimum values for its length and C. | A | T | :CALC[1-4]:SRL:CONN[1-2]:IMM | :CALC[1-4]:SRL:CONN[1-2]:IMM |  | Need to define the unit with METers or FEET with ":CALC:TRAN:DIST:UNIT" first
SENSe[1¦2]:CORRection:OFFSet:PHASe | Specifies the phase offset. | C | T | :CALC[1-4]:CORR:OFFS:PHAS | :CALC[1-4]:CORR:OFFS:PHAS |  | Eligible for active trace only
SENSe[1¦2]:CORRection:ONEPort:REFLection[:IMMediate] | Retrieves user one-port reflection correction factors from memory. | D | F
SENSe[1¦2]:CORRection:ONEPort:TRANSmission[:IMMediate] | Retrieves user one-port transmission correction factors from memory. | D | F
SENSe[1¦2]:CORRection:TWOPort[:IMMediate] | Retrieves user two-port correction factors from memory. | D | F
SENSe[1¦2]:CORRection:PEAK:COAX[:STATe]<ON¦OFF> | Turns multi-peak correction on or off | A | T | :CALC{1-4}:MARK:FUNC:MULT:TYPE PEAK | :CALC{1-4}:MARK:FUNC:MULT:TYPE PEAK
SENSe[1¦2]:CORRection:RVELocity:COAX <num> | Specifies the velocity factor. | A+ | T | :SENS[1-4]:CORR:RVEL:COAX | :SENS[1-4]:CORR:RVEL:COAX
SENSe[1¦2]:CORRection:RVELocity[:IMMediate] | Measures cable and determine optimum values for cable loss and velocity factor. | D | F
SENSe[1¦2]:CORRection:TESTSET | Brings up the Test Set Cal menu. | D | F |  |  |  | Use direct control commands
SENSe[1¦2]:CORRection:THReshold:COAX <num> | Selects multi-peak threshold value, in dB. | A | T | :CALC{1-4}:MARK:FUNC:MULT:PEXC <num> | :CALC{1-4}:MARK:FUNC:MULT:PEXC <num>
SENSe[1¦2]:COUPle <char> | Turns the alternate sweep mode on/off. | B | T | :DISP:SPL | :DISP:SPL | Ch_trace | Use one channel for OFF, 2 channels for ON
SENSe[1¦2]:DETector[:FUNCtion] <char> | Specifies which detection mode is used | D | F |  |  |  | Only narrowband is available. Broadband isn't supported
SENSe[1¦2]:DISTance:CENTer <num> | Sets the center distance for a fault location measurement | A | T | :CALC{1-4}:TRAN:DIST:CENT <num> | :CALC{1-4}:TRAN:DIST:CENT <num>
SENSe[1¦2]:DISTance:STARt <num> | Sets the start distance for a fault location measurement | A | T | :CALC{1-4}:TRAN:DIST:STAR <num> | :CALC{1-4}:TRAN:DIST:STAR <num>
SENSe[1¦2]:DISTance:UNITs <char> | Specifies distance units. | A | T | :CALC{1-4}:TRAN:DIST:UNIT ｛MET¦FEET} | :CALC{1-4}:TRAN:DIST:UNIT ｛MET¦FEET}
SENSe[1¦2]:DISTance:STOP <num> | Sets the stop distance for a fault location measurement | A | T | :CALC{1-4}:TRAN:DIST:STOP <num> | :CALC{1-4}:TRAN:DIST:STOP <num>
SENSe[1¦2]:FREQuency:CENTer <num> | Sets the center frequency of the RF source. | A+ | T | :SENS[1-4]:FREQ:CENT | :SENS[1-4]:FREQ:CENT
SENSe[1¦2]:FREQuency:MODE <char> | Sets the fault location meas to CENTer (bandpass) or LOWPass. | A | T | :CALC{1-4}:TRAN:DIST {BPAS¦LPAS} | :CALC{1-4}:TRAN:DIST {BPAS¦LPAS} |  | BPASs (band pass) is default
SENSe[1¦2]:FREQuency:SPAN <num> | Sets the frequency span of the RF source. | A+ | T | :SENS[1-4]:FREQ:SPAN | :SENS[1-4]:FREQ:SPAN
SENSe[1¦2]:FREQuency:SPAN:MAXimum <num> | Sets the max freq span of the RF source for bandpass fault location meas. | B | T | :CALC{1-4}:TRAN:DIST:SPAN <num> | :CALC{1-4}:TRAN:DIST:SPAN <num> |  | When number range beyond the maximum limit, span is set as maximum in its current setting.
SENSe[1¦2]:FREQuency:STARt <num> | Sets the start frequency of the RF source. | A+ | T | :SENS[1-4]:FREQ:STAR | :SENS[1-4]:FREQ:STAR
SENSe[1¦2]:FREQuency:STOP <num> | Sets the stop frequency of the RF source. | A+ | T | :SENS[1-4]:FREQ:STOP | :SENS[1-4]:FREQ:STOP
SENSe[1¦2]:FREQuency:ZSTOp <num> | Sets the Z cutoff frequency for cable impedance calculations. | B | T | :CALC{1-4}:SRL:AUTO ON :CALC{1-4}:SRL:IMPAUTO:CUT <num> | :CALC{1-4}:SRL:AUTO ON :CALC{1-4}:SRL:IMPAUTO:CUT <num>
SENSe[1¦2]:FUNCtion? | Queries the measurement funciton | C | T | :CALCulate[1-4]:PARameter[1-4]:DEFine ? | :CALCulate[1-4]:PARameter[1-4]:DEFine ? |  | Reply s-parameter format
SENSe[1¦2]:FUNCtion 'FLOC <num>,<num>' | Specifies receiver will meas ratio of the power (fault location). | B | T | :CALC{1-4}:TRAN:DIST:STAT :CALC{1-4}:TRAN:TIME:STAT | :CALC{1-4}:TRAN:DIST:STAT :CALC{1-4}:TRAN:TIME:STAT
SENSe[1¦2]:FUNCtion 'XFRequency:POWer <num>' | Specifies receiver will measure power into a specific detector. | D | F
SENSe[1¦2]:FUNCtion 'XFRequency:GDELay:RATio <num>,<num>' | Specifies receiver will measure ratio of the power ( group delay). | B | T | :CALC{1-4}:FORM GDEL :CALC{1-4}:PAR{1-4}:DEF S21 | :CALC{1-4}:FORM GDEL :CALC{1-4}:PAR{1-4}:DEF S22
SENSe[1¦2]:FUNCtion 'XFRequency:POWer:RATio <num>,<num>' | Specifies receiver will measure a ratio of the power. | C | T | :CALC[1-4]:PAR[1-4]:DEF | :CALC[1-4]:PAR[1-4]:DEF |  | Set with s-parameter definition. A/R=S11, B/R=S21
SENSe[1¦2]:FUNCtion 'XFRequency:S <num>,<num>' | Specifies receiver will measure an s-parameter. | C | T | :CALCulate[1-4]:PARameter[1-4]:DEFine ? | :CALCulate[1-4]:PARameter[1-4]:DEFine ?
SENSe[1¦2]:FUNCtion:FAULt:CONNector [:IMMediate] | Forces a connector verification measurement on the alternate ch. | C | T | :CALC{1-4}:SRL:CONN{1-2}:IMM | :CALC{1-4}:SRL:CONN{1-2}:IMM
SENSe[1¦2]:FUNCtion:SRL:<num>,<num> | Specifies receiver will measure ratio of the power (SRL) | A | T | :CALC{1-4}:SRL <bool> | :CALC{1-4}:SRL <bool>
SENSe[1¦2]:FUNCtion:SRL:IMPedance <num> | Sets the cable impedance. | A | T | :CALC[1-4]:SRL:MAN | :CALC[1-4]:SRL:MAN
SENSe[1¦2]:FUNCtion:SRL:MODE <char> | Sets the auto z function to AUTO or MANual. | A | T | :CALC[1-4]:SRL:AUTO | :CALC[1-4]:SRL:AUTO |  | When AUTO, select ON When Manual, select OFF
SENSe[1¦2]:FUNCtion:SRL:SCAN[:IMMediate] | Starts a cable scan. | B | T- | :MMEM:LOAD:PROG "XXX/XXX.vba" :PROG:NAME ""YYY.main"" :PROG:STAT RUN | :MMEM:LOAD:PROG "XXX/XXX.vba" :PROG:NAME ""YYY.main"" :PROG:STAT RUN |  | XXX/ YYY is defined later
SENSe[1¦2]:ROSCillator:SOURce <char> | Specifies the source of the reference oscillator, INT or EXT. | D | F | :SENSe[1-4]:ROSCillator:SOURce? | :SENSe[1-4]:ROSCillator:SOURce? |  | Query only
SENSe[1¦2]:STATe <ON¦OFF> | Turns the specified channel on/off . | C | T | :DISP:SPL | :DISP:SPL |  | Channels are always on. Controlled by display and trigger commands.
SENSe[1¦2]:SWEep:POINts <num> | Sets the number of data points for the measurement | A+ | T | :SENSe[1-4]:SWEep:POINts | :SENSe[1-4]:SWEep:POINts
SENSe[1¦2]:SWEep:TIME <num> | Sets the sweep time. | A+ | T | :SENSe[1-4]:SWEep:TIME[:DATA] | :SENSe[1-4]:SWEep:TIME[:DATA]
SENSe[1¦2]:SWEep:TIME:AUTO <ON¦OFF¦ONCE> | Turns the automatic sweep time function on/off. | A+ | T | :SENSe[1-4]:SWEep:TIME:AUTO | :SENSe[1-4]:SWEep:TIME:AUTO
SENSe:SWEep:TRIGger:SOURce <char> | Sets the trigger source for each point in a sweep, IMM or EXT. | D | F |  |  |  | Point trigger is not supported
SENSe[1¦2]:WINDow[:TYPE] | Sets the window selection for fault location measurements. | C | T | :CALC{1-4}:TRAN:DIST:KBES <num> :CALC{1-4}:TRAN:TIME:KBES <num> :CALC{1-4}:TRAN:DIST:IMP:WIDT <num> :CALC{1-4}:TRAN:TIME:IMP:WIDT <num> :CALC{1-4}:TRAN:DIST:STEP:RTIM <num> :CALC{1-4}:TRAN:TIME:STEP:RTIM <num> | :CALC{1-4}:TRAN:DIST:KBES <num> :CALC{1-4}:TRAN:TIME:KBES <num> :CALC{1-4}:TRAN:DIST:IMP:WIDT <num> :CALC{1-4}:TRAN:TIME:IMP:WIDT <num> :CALC{1-4}:TRAN:DIST:STEP:RTIM <num> :CALC{1-4}:TRAN:TIME:STEP:RTIM <num> |  | Soft key's max, min, and normal are correspond with 13, 0, and 6 respectively
SOURce[1¦2]:POWer[:LEVel][:IMMediate][:AMPLitude]<num> | Sets the RF power output from the source. | A+ | T | :SOURce[1-4]:POWer[:LEVel][:IMMediate][:AMPLitude] | :SOURce[1-4]:POWer[:LEVel][:IMMediate][:AMPLitude] |  | -5dBm to +10dBm @wo/opt.1E1 -45dBm to +10dBm @w/opt.1E1
SOURce:POWer:PRESet <num> | Sets the power level that NA will always return to after preset. | D | F |  |  |  | fixed to 0 dBm
SOURce[1¦2]:POWer:RANGe <char> | Specifies the power sweep range (ATT10 to ATT60). | A+ | T | :SOURce[1-4]:POWer:ATTenuation[:DATA] | :SOURce[1-4]:POWer:ATTenuation[:DATA] |  | define ATT value from 0 dB to 35 dB with 5 dB step
SOURce[1¦2]:POWer:STARt <num> | Sets the power sweep start power. | A+ | T | :SOURce[1-4]:POWer:STARt | :SOURce[1-4]:POWer:STARt
SOURce[1¦2]:POWer:STOP <num> | Sets the power sweep stop power. | A+ | T | :SOURce[1-4]:POWer:STOP | :SOURce[1-4]:POWer:STOP
STATus:DEVice:CONDition? | Reads the Device Status condition register. | D | F |  |  | STB | Device Status Resister is not supported
STATus:DEVice:ENABle <num> | Sets and queries bits in the Device Status enable register. | D | F |  |  | STB | Device Status Resister is not supported
STATus:DEVice[:EVENt]? | Reads and clears the Device Status event register. | D | F |  |  | STB | Device Status Resister is not supported
STATus:DEVice:NTRansition <num> | Sets and queries bits in the Device Status negative transition register. | D | F |  |  | STB | Device Status Resister is not supported
STATus:DEVice:PTRansition <num> | Sets and queries bits in the Device Status positive transition register. | D | F |  |  | STB | Device Status Resister is not supported
STATus:OPERation:AVERaging:CONDition? | Reads the Averaging status condition register. | D | F |  |  | STB | Operation Resister is not enough to read out
STATus:OPERation:AVERaging:ENABle <num> | Sets and queries bits in the Averaging status enable register. | D | F |  |  | STB | Operation Resister is not enough to read out
STATus:OPERation:AVERaging[:EVENt]? | Reads and clears the Averaging status event register. | D | F |  |  | STB | Operation Resister is not enough to read out
STATus:OPERation:AVERaging:NTRansition <num> | Sets and queries bits in the Averaging status negative transition register. | D | F |  |  | STB | Operation Resister is not enough to read out
STATus:OPERation:AVERaging:PTRansition <num> | Sets and queries bits in the Averaging status positive transition register. | D | F |  |  | STB | Operation Resister is not enough to read out
STATus:OPERation:CONDition? | Reads the Operational status condition register. | C | T | :STATus:OPERation:CONDition? | :STATus:OPERation:CONDition? | STB
STATus:OPERation:ENABle <num> | Sets and queries bits in the Operational status enable register. | C | T | :STATus:OPERation:ENABle | :STATus:OPERation:ENABle | STB
STATus:OPERation[:EVENt]? | Reads and clears the Operational status event register. | C | T | :STATus:OPERation[:EVENt]? | :STATus:OPERation[:EVENt]? | STB
STATus:OPERation:MEASuring:CONDition? | Reads the Measuring status condition register. | C | T | :STATus:OPERation:CONDition? | :STATus:OPERation:CONDition? |  | Bit 4 (16) express all channels' sweep state only
STATus:OPERation:MEASuring:ENABle <num> | Sets and queries bits in the Measuring status enable register. | C | T | :STATus:OPERation:ENABle 4 | :STATus:OPERation:ENABle 5 |  | Bit 4 (16) express all channels' sweep state only
STATus:OPERation:MEASuring[:EVENt]? | Reads and clears the Measuring status event register. | C | T | :STATus:OPERation[:EVENt]? | :STATus:OPERation[:EVENt]? |  | Bit 4 (16) express all channels' sweep state only
STATus:OPERation:MEASuring:NTRansition <num> | Sets and queries bits in the Measuring status negative transition register. | C | T | :STATus:OPERation:NTRansition | :STATus:OPERation:NTRansition |  | Bit 4 (16) express all channels' sweep state only
STATus:OPERation:MEASuring:PTRansition <num> | Sets and queries bits in the Measuring status positive transition register. | C | T | :STATus:OPERation:PTRansition | :STATus:OPERation:PTRansition |  | Bit 4 (16) express all channels' sweep state only
STATus:OPERation:NTRansition <num> | Sets and queries bits in the Operational Status negative transition register. | C | T | :STATus:OPERation:NTRansition | :STATus:OPERation:NTRansition
STATus:OPERation:PTRansition <num> | Sets and queries bits in the Operational Status positive transition register. | C | T | :STATus:OPERation:PTRansition | :STATus:OPERation:PTRansition
STATus:PRESet | Sets bits in most enable and transition registers to their default state. | A+ | T | :STATus:PRESet | :STATus:PRESet
STATus:QUEStionable:CONDition? | Reads the Questionable Status condition register. | C | T | :STATus:QUEStionable:CONDition? | :STATus:QUEStionable:CONDition? | STB | Data Questionable is not supported, though Limit fail is supported
STATus:QUEStionable:ENABle <num> | Sets and queries bits in the Questionable Status enable register. | C | T | :STATus:QUEStionable:ENABle | :STATus:QUEStionable:ENABle | STB | Data Questionable is not supported, though Limit fail is supported
STATus:QUEStionable[:EVENt]? | Reads and clears the Questionable Status event register. | C | T | :STATus:QUEStionable[:EVENt]? | :STATus:QUEStionable[:EVENt]? | STB | Data Questionable is not supported, though Limit fail is supported
STATus:QUEStionable:LIMit:CONDition? | Reads the Limit Fail condition register. | C | T | :STATus:QUEStionable:LIMit:CONDition? | :STATus:QUEStionable:LIMit:CONDition? | STB | Marker Limit is not supported, though Trace limit is supported
STATus:QUEStionable:LIMit:ENABle <num> | Sets and queries bits in the Limit Fail enable register. | C | T | :STATus:QUEStionable:LIMit:ENABle | :STATus:QUEStionable:LIMit:ENABle | STB | Marker Limit is not supported, though Trace limit is supported
STATus:QUEStionable:LIMit[:EVENt]? | Reads and clears the Limit Fail event register. | C | T | :STATus:QUEStionable:LIMit[:EVENt]? | :STATus:QUEStionable:LIMit[:EVENt]? | STB | Marker Limit is not supported, though Trace limit is supported
STATus:QUEStionable:LIMit:NTRansition <num> | Sets and queries bits in the Limit Fail negative transition register. | C | T | :STATus:QUEStionable:LIMit:NTRansition | :STATus:QUEStionable:LIMit:NTRansition | STB | Marker Limit is not supported, though Trace limit is supported
STATus:QUEStionable:LIMit:PTRansition <num> | Sets and queries bits in the Limit Fail positive transition register. | C | T | :STATus:QUEStionable:LIMit:PTRansition | :STATus:QUEStionable:LIMit:PTRansition | STB | Marker Limit is not supported, though Trace limit is supported
STATus:QUEStionable:NTRansition <num> | Sets and queries bits in the Questionable Status negative transition register. | C | T | :STATus:QUEStionable:LIMit:NTRansition | :STATus:QUEStionable:LIMit:NTRansition | STB | Data Questionable is not supported, though Limit fail is supported
STATus:QUEStionable:PTRansition <num> | Sets and queries bits in the Questionable Status positive transition register. | C | T | :STATus:QUEStionable:LIMit:PTRansition | :STATus:QUEStionable:LIMit:PTRansition | STB | Data Questionable is not supported, though Limit fail is supported
SYSTem:BEEPer[:IMMediate][<freq>[,<dur>[,<vol>] | Instructs the analyzer to beep. | C | T | :SYSTem:BEEPer:COMPlete:IMMediate | :SYSTem:BEEPer:COMPlete:IMMediate | STB | Cannot control parameters such as freq, dur, and vol.
SYSTem:BEEPer:VOLume <num> | Sets the volume of the beeper | D | F
SYSTem:COMMunicate:GPIB:CONTroller[:STATe]<ON¦OFF> | Makes the analyzer the system controller. | D | F
SYSTem:COMMunicate:GPIB:ECHO <ON¦OFF> | Turns HP-IB mnemonic echo on/off . | D | F
SYSTem:COMMunicate:GPIB:HCOPy:ADDRess <num> | Sets the address of an HP-IB printer or plotter for hardcopy output | D | F
SYSTem:COMMunicate:GPIB[:SELF]:ADDRess <num> | Sets the analyzer's HP-IB address | D | F
SYSTem:COMMunicate:LAN:BOOTp:HOST <string> | Sets the host name or host IP address of the remote host. | D | F
SYSTem:COMMunicate:LAN:BOOTp:STATE <OFF¦0¦ON¦1> | Sets a BOOTP request when the analyzer boots up. | D | F
SYSTem:COMMunicate:LAN:BOOTp:TIMeout#1~MAX_AUTO_CAL_TIME[S]# | Sets the number of seconds the NA will retry the BOOTP requests. | D | F
SYSTem:COMMunicate:LAN:BOOTp:TRANsfer:FILE:NAME <string> | Sets the path file name of the boot file you want to receive at boot time. | D | F
SYSTem:COMMunicate:LAN:BOOTp:TRANsfer:FTP:PASSword-? <string> | Selects the password of the remote BOOTP host. | D | F
SYSTem:COMMunicate:LAN:BOOTp:TRANsfer:FTP:USERname <string> | Selects the user name of the remote BOOTP host. | D | F
SYSTem:COMMunicate:LAN:BOOTp:TRANsfer:METHod <TFTP¦FTP> | Selects TFTP as the file transfer program. | D | F
SYSTem:COMMunicate:LAN:EADDress? | Queries the analyzer's ethernet address. | D | F |  |  |  | Only front key access
SYSTem:COMMunicate:LAN:IPADdress <string> | Sets the analyzer's Internet Protocol address. | D | F |  |  |  | Only front key access
SYSTem:COMMunicate:LAN:LOGin:USER:ADD-? <string>,<string> | Selects the login user/password pairs. | D | F
SYSTem:COMMunicate:LAN:LOGin:USER:DELETE-? <string>,<string> | Deletes the login user/password pairs. | D | F
SYSTem:COMMunicate:LAN:LOGin:USER:LIST:COUNt? | Queries login user names. | D | F
SYSTem:COMMunicate:LAN:LOGin:USER:LIST:NAME? #1-7 | Queries user name. | D | F
SYSTem:COMMunicate:LAN:NFS:AUTHentiation:ID:GROup#0~4.74836e+07# | Selects the NFS remote file system for entering a Group ID. | D | F
SYSTem:COMMunicate:LAN:NFS:AUTHentiation:ID:USER#0~4.74836e+07# | Selects the NFS remote file system for entering a User ID. | D | F
SYSTem:COMMunicate:LAN:NFS:AUTOmount:ADD-? <string>,<string>,<string> | Sets the mounted NFS device to the automount device table. | D | F
SYSTem:COMMunicate:LAN:NFS:AUTOmount:LIST:COUNt? | Selects the number of NFS devices in Automount list. | D | F
SYSTem:COMMunicate:LAN:NFS:AUTOmount:LIST:LOCFilesys? #1-7 | Selects the local file system name for an NFS device. | D | F
SYSTem:COMMunicate:LAN:NFS:AUTOmount:LIST:REMFilesys? #1-7 | Queries the automount list. | D | F
SYSTem:COMMunicate:LAN:NFS:AUTOmount:LIST:REMHost? #1-7 | Queries the remote host name (or host IP address) for NFS device. | D | F
SYSTem:COMMunicate:LAN:NFS:AUTOmount:REMove-? <string> | Removes device from automount list. | D | F
SYSTem:COMMunicate:LAN:NFS:MOUNT-? <string>,<string>,<string> | Selects remote host, file system, and local system name for NFS device. | D | F
SYSTem:COMMunicate:LAN:NFS:MOUNT:LIST:COUNt? | Queries the number of currently mounted NFS devices. | D | F
SYSTem:COMMunicate:LAN:NFS:MOUNT:LIST:LOCFilesys? #1-7 | Queries the local file system name. | D | F
SYSTem:COMMunicate:LAN:NFS:MOUNT:LIST:REMFilesys? #1-7 | Selects the remote host file system name for NFS device. | D | F
SYSTem:COMMunicate:LAN:NFS:MOUNT:LIST:REMHost? #1-7 | Selects the remote host name (or host IP address) for NFS device. | D | F
SYSTem:COMMunicate:LAN:NFS:UNMount-? <string> | Selects the NFS device table. | D | F
SYSTem:COMMunicate:LAN:PRINter:HOSTname <string> | Specifies the IP address of the LAN printer. | D | F
SYSTem:COMMunicate:LAN:ROUTe:GATeway <string> | Sets the IP address for a LAN gateway. | D | F
SYSTem:COMMunicate:LAN:ROUTe:SMASk <string> | Sets the subnet mask. | D | F |  |  |  | Only front key access
SYSTem:COMMunicate:LAN:SCPI:SOCKet:PORT <num> | Selects port number for socket connection to NA for SCPI socket programming. | D | F
SYSTem:COMMunicate:LAN:SICL:HPIB:ADDRess #0~20, 22~30# | Sets the HPIB address for SICL LAN. | D | F
SYSTem:COMMunicate:LAN:SICL:HPIB:LU #0~1024# | Sets the HPIB logical unit number for SICL LAN. | D | F
SYSTem:COMMunicate:LAN:SICL:HPIB:NAME <string> | Sets the HPIB name for SICL LAN. | D | F
SYSTem:COMMunicate:SERial:TRANsmit:BAUD <num> | Sets the baud rate for hardcopy output to a device on the serial port | D | F
SYSTem:COMMunicate:SERial:TRANsmit:HANDshake <char> | Sets handshake for communication to hardcopy device on the serial port. | D | F
SYSTem:COMMunicate:TTL:USER:FEED <char> | Selects the function of the USER TTL IN/OUT port on the rear panel. | D | F |  |  |  | USER TTL is not supported
SYSTem:DATE <num1>,<num2>,<num3> | Sets the year (num1), month (num2) and day (num3) of the real time clock. | A+ | T | :SYSTem:DATE | :SYSTem:DATE
SYSTem:ERRor? | Queries the error queue. | A+ | T | :SYSTem:ERRor? | :SYSTem:ERRor?
SYSTem:KEY <char> | Sends key names which execute the same functions as front panel keys. | D | F
SYSTem:KEY:MASK? | Queries the mask (shift, ctrl, alt) associated with a keypress on kbd. | D | F
SYSTem:KEY:QUEue:CLEar | Clears the key queue. | D | F
SYSTem:KEY:QUEue:COUNt? | Queries the number of key codes in the queue. | D | F
SYSTem:KEY:QUEue:MAXimum? | Queries the size of the key queue. | D | F
SYSTem:KEY:QUEue[:STATe] <ON¦OFF> | Turns on/off the key queue. | D | F
SYSTem:KEY:TYPE? | Queries the type of key that was pressed. | D | F
SYSTem:KEY:USER | Sets the User Request bit of the Standard Event Status Register. | D | F
SYSTem:KEY[:VALue]? | Queries the key code value for the last key pressed | D | F
SYSTem:PRESet | Performs a system preset. | A+ | T
SYSTem:SET <block> | Sends a learn string to the analyzer | D | F
SYSTem:SET:LRN? [<USER>] | Queries or set the instrument state | D | F
SYSTem:SET:LRNLong? [<USER>] | Queries or set the instrument state, data, and calibration. | D | F
SYSTem:TIME <num1>,<num2>,<num3> | Sets the hour, minute and second of the real time clock. | A+ | T | :SYSTem:TIME | :SYSTem:TIME
SYSTem:VERSion? | Queries the SCPI version of the analyzer. | D | F
TEST:RESult? | Queries the result of the selected adjustment or self-test | D | F |  |  |  | No commands for self diagnostics
TEST:SELect <num> | Selects the adjustment or self-test to execute. | D | F
TEST:STATe <char> | Selects the state of the active adjustment or self-test | D | F
TEST:VALue <num> | Sets or queries a value for an adjustment or self-test. | D | F
TRACe[:DATA]? <char> | Queries trace and calibration coefficient data. CH<1¦2>FDATA, CH<1¦2>FMEM, CH<1¦2>SDATA, CH<1¦2>SMEM, CH<1¦2><A¦B¦R>FWD, CH<1¦2>SCORR<X> | B | T | :CALCulate[1-4][:SELected]:DATA:FDATa :CALCulate[1-4][:SELected]:DATA:FMEMory :CALCulate[1-4][:SELected]:DATA:SDATa :CALCulate[1-4][:SELected]:DATA:SMEMory :SENSe[1-4]:CORRection:COEFficient? <enum>, <rcv_port>, <src_port> | :CALCulate[1-4][:SELected]:DATA:FDATa :CALCulate[1-4][:SELected]:DATA:FMEMory :CALCulate[1-4][:SELected]:DATA:SDATa :CALCulate[1-4][:SELected]:DATA:SMEMory :SENSe[1-4]:CORRection:COEFficient? <enum>, <rcv_port>, <src_port> | Data | FWD is not supported <enum> = {ES¦ER¦ED¦EL¦ET¦EX}
TRACe[:DATA] <char>,<data> | Inputs trace data and calibration coefficient data. | B | T | :CALCulate[1-4][:SELected]:DATA:FDATa :CALCulate[1-4][:SELected]:DATA:FMEMory :CALCulate[1-4][:SELected]:DATA:SDATa :CALCulate[1-4][:SELected]:DATA:SMEMory | :CALCulate[1-4][:SELected]:DATA:FDATa :CALCulate[1-4][:SELected]:DATA:FMEMory :CALCulate[1-4][:SELected]:DATA:SDATa :CALCulate[1-4][:SELected]:DATA:SMEMory | Data | FWD and SCORR are not supported
TRACe:CORRection:SIMulate[:DATA] {STD1¦STD2¦...¦STD12} | Reads simulated calibration standards. | D | F
TRACe:CORRection:SIMulate SAVE <TRAN1¦TRAN2¦TRAN3¦REFL3¦TESTset¦TESTset1¦TESTset2¦VERIFY¦¦TWOPort¦NONE> | Computes cal error correction arrays. | D | F
TRIGger[:SEQuence]:SOURce <char> | Sets the source for the sweep trigger signal | A+ | T | :TRIGger[:SEQuence]:SOURce | :TRIGger[:SEQuence]:SOURce
*CLS |  | A+ | T | *CLS | *CLS
*ESE |  | C | T | *ESE | *ESE | STB
*ESE? |  | C | T | *ESE? | *ESE? | STB
*ESR? |  | C | T | *ESR? | *ESR? | STB
*IDN? |  | A+ | T | *IDN? | *IDN?
*LRN? |  | D | F
*OPC |  | A+ | T | *OPC | *OPC
*OPC? |  | A+ | T | *OPC? | *OPC?
*OPT? |  | C | T | *OPT? | *OPT?
*PCB |  | D | F
*PSC |  | D | F
*RST |  | C | T | *RST | *RST
*SRE |  | C | T | *SRE | *SRE | STB
*SRE? |  | C | T | *SRE? | *SRE? | STB
*STB? |  | C | T | *STB? | *STB? | STB
*TRG |  | C | T | *TRG :TRIG :TRIG:SING | *TRG :TRIG :TRIG:SING | Trigger
*TST? |  | D | F
*WAI |  | A+ | T | *WAI | *WAI
ABORt | Aborts and resets the sweep in progress. | A+ | T | ABORt | ABORt